Supplementary table S1: Summary of a representative acquisition method for BSA, for Orbitrap Fusion Tribrid (Thermo Scientific) with Instrument Control Software 3.0

Application Mode:   Peptide

Method Duration (min):   30

Global Parameters

Ion Source

Ion Source Type:   NSI

Spray Voltage:   Static

Positive Ion (V):   0

Negative Ion (V):   600

Positive Ion

| Positive Ion |  |
| --- | --- |
| Time (min) | Voltage (V) |
| 0 | 0 |

Negative Ion

| Negative Ion |  |
| --- | --- |
| Time (min) | Voltage (V) |
| 0 | 0 |

Sweep Gas (Arb):   0

Ion Transfer Tube Temp (°C):   170

Use Ion Source Settings from Tune:   False

FAIMS Mode:   Not Installed

MS Global Settings

Infusion Mode:   Liquid Chromatography

Expected LC Peak Width (s):   30

Default Charge State:   1

Internal Mass Calibration:   Off

Start Time (min):   0

End Time (min):   30

Cycle Time (sec):   50000

Master Scan:

MS OT

Detector Type:   Orbitrap

Orbitrap Resolution:   240000

Mass Range:   Normal

Use Quadrupole Isolation:   True

Scan Range (m/z):   200-1200

RF Lens (%):   65

Master Scan: MS OT (Contd)

AGC Target:   Custom

Normalized AGC Target (%):   1250

Maximum Injection Time Mode:   Custom

Maximum Injection Time (ms):   3000

Microscans:   10

Data Type:   Profile

Polarity:   Positive

Source Fragmentation:   Disabled

Scan Description:

Filters:

MIPS

Monoisotopic Peak Determination:   Peptide

Data Dependent

Data Dependent Mode:   Cycle Time

Time between Master Scans (sec):   50000

Scan Event Type 1:

Scan:

ddMS² OT HCD

Isolation Mode:   Quadrupole

Isolation Window (m/z):   0.4

Isolation Offset:   Off

Activation Type:   HCD

Collision Energy Mode:   Fixed

HCD Collision Energy (%):   30

Detector Type:   Orbitrap

Orbitrap Resolution:   7500

Mass Range:   Normal

Scan Range Mode:   Auto

AGC Target:   Custom

Normalized AGC Target (%):   40

Maximum Injection Time Mode:   Custom

Maximum Injection Time (ms):   3000

Microscans:   1

Data Type:   Centroid

Use EASY-IC™:   False

Scan Description: