Appendix

The PulseSPEL program for the PELDOR experiment is given below. In order to use it, copy the code into a plain text file and save it under the file name “PELDOR.exp”.

; Double Electron-Electron Resonance

; PulseSPEL pulse programs, phase programs & dimensions

; Dimensions

; -------------------------------------------

begin defs          ; dimension of data array [sx,sy]

dim s[256]					     											    ; for 2P ESE Setup

dim1 s[256]        ; for 2P Field Sweep

dim2 s[256]        ; for 3P ELDOR Setup

dim3 s[100]        ; for 3P ELDOR Nutation

dim4 s[256]        ; for 4P DEER Setup

dim5 s[379]        ; for 4P DEER sx=(d1+d2-2*d3)/d30

end defs

; Phase cycling programs

; -------------------------------------------

begin lists "+<x> none"

ph1 +<x>

ph2 +<x>

asg1 +a

bsg1 +b

end lists

begin lists1 "-<x> none"

ph1 -<x>

ph2 -<x>

asg1 +a

bsg1 +b

end lists1

begin lists2 "2-step"

ph1 +<x> -<x>

ph2 +<x>

asg1 +a -a

bsg1 +b -b

end lists2

; PulseSPEL Programs

; Standing 2-pulse Hahn echo

; -------------------------------------------

begin exp "2P ESE Setup" [TRANS QUAD]

shot i=1 to h           ; accumulation loop

d9                   ; DAF

p0 [ph1]              ; 1st pulse

d1                   ; tau1

p1 [ph2]              ; 2nd pulse

d1                   ; tau1

d0                   ; acquistion delay

dig [sg1]              ; acquisition

next i                  ; end of accumulation loop

end exp

; 2-pulse Field Sweep

; -------------------------------------------

begin exp1 "2P Field Sweep" [INTG QUAD]

for k=1 to n            ; averaging loop

totscans(n)           ; output of total number of scans

bsweep x=1 to sx													     ; field sweep loop

shot i=1 to h        ; accumulation loop

d9                ; DAF

p0 [ph1]           ; 1st pulse

d1                ; tau1

p1 [ph2]           ; 2nd pulse

d1                ; tau1

d0                ; acquistion delay

acq [sg1]          ; acquisition

next i              ; end of accumulation loop

next x               ; end of field sweep loop

scansdone(k)        ; output of scans done

next k                ; end of averaging loop

end exp1

; Standing 3-pulse Echo

; -------------------------------------------

begin exp2 "3P ELDOR Setup" [TRANS QUAD]

shot i=1 to h           ; accumulation loop

d9                   ; DAF

p2 [ELDOR]          ; 1st pulse

d4                   ; T

p0 [ph1]              ; 2nd pulse

d1                   ; tau1

p1 [ph2]              ; 3d pulse

d1                   ; tau1

d0                   ; acquistion delay

dig [sg1]              ; acquisition

next i                  ; end of accumulation loop

end exp2

; Pump Pulse Optimization

; -------------------------------------------

begin exp3 "3P ELDOR Nutation" [SPT QUAD]

for k=1 to n            ; averaging loop

totscans(n)           ; output of total number of scans

dx=p3                ; define time axis

p4=p3                ; initial pump pulse length

d5=d4                ; initial T

sweep x=1 to sx       ; sweep loop

shot i=1 to h         ; accumulation loop

d9                ; DAF

p4 [ELDOR]        ; 1st pulse

d5                 ; T

p0 [ph1]            ; 2nd pulse

d1                 ; tau1

p1 [ph2]            ; 3d pulse

d1                 ; tau1

d0                 ; acquistion delay

acq [sg1]           ; acquisition

next i               ; end of accumulation loop

dx=dx+d6           ; define time axis increment

p4=p4+d6           ; increment pump pulse length

d5=d5+d6           ; increment T

next x                ; end of sweep loop

scansdone(k)         ; output of scans done

next k                 ; end of averaging loop

end exp3

; Standing 4-pulse DEER

; -------------------------------------------

begin exp4 "4P DEER Setup" [TRANS QUAD]

shot i=1 to h           ; accumulation loop

d9                  ; DAF

p0 [ph1]             ; 1st pulse

d1                  ; tau1

p1 [ph2]             ; 2nd pulse

d1                  ; tau1

p2 [ELDOR]         ; 1st pulse

d2                  ; tau2

p1 [ph2]             ; 3d pulse

d2                  ; tau2

d0                  ; acquistion delay

dig [sg1]             ; acquisition

next i                 ; end of accumulation loop

end exp4

; 4-pulse DEER

; -------------------------------------------

begin exp5 "4P DEER" [INTG QUAD]

for k=1 to n            ; averaging loop

totscans(n)           ; output of total number of scans

d10=d1              ; initial tau1

d13=d3              ; initial deadtime delay

for j=1 to m           ; tau1 averaging loop

dx=0                ; define time axis

d11=d13            ; tau1 + initial t

d12=d1+d2          ;

d12=d12-d3         ; tau2 - initial t

sweep x=1 to sx      ; sweep loop

shot i=1 to h        ; accumulation loop

d9                ; DAF

p0 [ph1]           ; 1st pulse

d10               ; tau1

p1 [ph2]           ; 2nd pulse

d11               ; tau1

p2 [ELDOR]        ; 1st pulse

d12               ; tau2

p1 [ph2]           ; 3d pulse

d2                ; tau2

d0                ; acquistion delay

acq [sg1]          ; acquisition

next i              ; end of accumulation loop

dx=dx+d30         ; define time axis increment

d11=d11+d30       ; tau1 + t

d12=d12-d30       ; tau2 - t

next x                ; end of sweep loop

d10=d10+d31         ; increment tau1

d13=d13+d31         ; increment initial deadtime delay

next j                 ; end of tau1 averaging loop

scansdone(k)          ; output of scans done

next k                  ; end of averaging loop

end exp5

An example of the variable definitions is given below. In order to use it, copy the code into a plain text file and save it under the file name “PELDOR_variable_definitions.def”.

; Double Electron-Electron Resonance

; PulseSPEL general variables, definitions & conventions

;

begin defs

; Variables    ; Comments

p0   = 12      ; 90 pulse length, pi/2

p1   = 24      ; 180 pulse length, pi

p2   = 16      ; pump pulse length

p3   = 2       ; the shortest accessible p2

p4   = 2       ; same as p3 (superseeded)

pg   = 24      ; Integrator Gate Width

d0   = 360     ; Detection Trigger offset time

d1   = 232     ; tau1

d2   = 3000    ; tau2

d3   = 100     ; deadtime delay between probe pi-pulse and first position of pump pulse

d4   = 400     ; T for pump pulse optimization

d5   = 500     ; same as d4 (superseeded)

d6   = 2       ; increment for pump pulse length

d9   = 0       ; DAF (Delay After Flash)

d10  = 232     ; same as tau1 (superseeded)

d11  = 100     ; same as deadtime delay(superseeded)

d12  = 3232    ; same as (tau1 + tau2 - deadtime delay) (superseeded)

d13  = 100     ; same as deadtime delay (superseeded)

d30  = 8       ; 1st time scale increment

d31  = 16      ; 2nd time scale increment (for modulation averaging)

a    = 50      ; number of transient averages (TRANS)

h    = 10      ; number of shots per point (INTG and SPT)

n    = 1       ; number of sweeps to accumulate

m    = 8       ; number of tau1 increments (for modulation averaging)

srt  = 3000 * srtu ; (RESERVED) SRT - Shot Repetition Time (srtu=1.02 us)

b    = 1       ;

c    = 1       ;

r    = 1       ;

w    = 180    ; Split point of sweeps (points), i.e. pulse separation where TWT gate pulse can be split

s    = 300    ; Sweep length (n. of data really taken)   ( counter: X )

t    = 1       ; Second time axsis sweep length           ( counter: Y )

dx   = 0      ; 1st time scale starting value

dy   = 0      ; 2nd time scale starting value

end defs